ALLEN ORTHOTICS AND PROSTHETICS

2502 West Ohio Avenue					                   1921 East 37th Street, Suite A

Midland, Texas							                              Odessa, Texas

432-683-3788								                                432-332-9821

APPLICATION INSTRUCTIONS FOR BK PROSTHESIS

WITH LOCKING LINER

Application of the Prosthesis:

- When utilizing a BK prosthesis with a locking gel liner, the first thing placed over your residual limb will be the locking gel liner. This is applied directly over the skin to reduce rubbing and is considered the initial layer.  Be sure the pin is in line with your residual limb.

- A wool, cotton or synthetic prosthetic sock(s) is placed over the locking liner, being careful that no wrinkles are present in the sock(s).  Your prosthetist will help you to determine the correct number and ply of socks to be worn. If more than one prosthetic sock is necessary, pull each sock on individually. Also, apply the heaviest socks on your limb first (after you have applied the initial layer). For example, first apply the initial layer, then any five-ply socks, then any three-ply socks, and finally, any one-ply socks go on top. The heavier socks will not tend to wrinkle as readily as the thinner ones and this will help protect your skin.

- Visually inspect the inside of the socket and locking mechanism before applying the prosthesis, to insure that no debris or foreign objects have fallen inside the socket or locking mechanism.  Apply the prosthesis over the locking liner and socks, making sure the knee-cap, is centered in the front of the prosthesis.  Push your limb firmly into the socket engaging the pin into the lock mechanism.

**  Instructions For Different Locking Mechanisms:

Click Type- When pushing your limb into the socket you will hear an audible click which indicates that the attachment pin has engaged the locking mechanism. It is absolutely necessary that you hear at least one click; however, should you hear multiple clicks, that is acceptable.  To insure that the pin has engaged you should attempt to pull the prosthesis off your residual limb before standing.

Clutch/Smooth Pin Type - When pushing your limb into the socket you will NOT hear an audible click.  To insure that the pin has engaged you should attempt to pull the prosthesis off your residual limb before standing.

Gator Grip Type - When pushing your limb into the socket you will hear ONE audible click which indicates that the attachment pin has engaged the locking mechanism. It is absolutely necessary that you hear one click.  To insure that the pin has engaged you should attempt to pull the prosthesis off your residual limb before standing.

- When the limb is completely into the socket the Patella should rest in the U-shaped window in the front of the socket.  The bottom of the residual limb should be in slight contact with the bottom of the socket with all your weight on the prosthesis.  If the residual limb does not go all the way into the socket with full weight bearing check to make sure the sock is not wrinkled up inside the socket.

If the residual limb is still not in the proper position after weight shifting several times then reduce the sock thickness.  If the residual limb goes too far into the socket, increase the sock thickness.

- To remove the prosthesis, press the release button on the side of the prosthesis.  This will disengage the pin from the locking mechanism, allowing you to remove the residual limb from the prosthesis.

Report failures, malfunctions or changes in your physical condition to us as soon as possible.